## Slide 1: Multi-lens ultrasound arrays enable large scale three-dimensional micro-vascularization characterization over whole organs
Nabil Haidour , Hugues Favre , Philippe Mateo, Juliette Reydet , Alain Bizé , Lucien Sambin , Jianping Dai , Paul-Matthieu Chiaroni, Bijan Ghaleh , Mathieu Pernot , Mickael Tanter & Clement Papadacci
1 Institute Physics for Medicine Paris – INSERM, ESPCI Paris-PSL, CNRS, Paris, France.
2 INSERM U955-IMRB, University Paris-Est Creteil, National Veterinary School of Alfort, Paris, France.
3 Interventional Cardiology Department, Henri Mondor Hospital, Creteil, France. e-mail: clement.papadacci@espci.fr

## Slide 2
Tissue oxygenation & nutrient exchange
Optimal organ function (heart, kidney, liver)
Structural or functional alteration in these networks :
Microcirculation is critical for :
Clinical motivation
Heart failure
Renal failure
various chronic diseases
Early detection of :
Tumor angiogenesis
Renal failure
Ischemia

## Slide 3
Limitations of existing modalities
CT angiography
✅ Clinical gold standard
✅ Whole organ
✅ Detailed anatomical insights
❌ Static nature
❌ Radiation

## Slide 4
Limitations of existing modalities
CT angiography
✅ Clinical gold standard
✅ Whole organ
✅ Detailed anatomical insights
❌ Static nature
❌ Radiation
4D Flow MRI
✅ Whole-organ or whole-body visualization
✅ Dynamic
❌ Limited spatial resolution : tenths of a millimeter
❌ Acquisition time
❌ Expensive

## Slide 5
Limitations of existing modalities
CT angiography
✅ Clinical gold standard
✅ Whole organ
✅ Detailed anatomical insights
❌ Static nature
❌ Radiation
4D Flow MRI
✅ Whole-organ or whole-body visualization
✅ Dynamic
❌ Limited spatial resolution : tenths of a millimeter
❌ Acquisition time
❌ Expensive
Doppler Ultrasound imaging
✅ Real time
✅ Blood flow assessment
✅ No radiation
❌ Low resolution
❌ Mainly 2D

## Slide 6
Limitations of existing modalities
CT angiography
✅ Clinical gold standard
✅ Whole organ
✅ Detailed anatomical insights
❌ Static nature
❌ Radiation
4D Flow MRI
✅ Whole-organ or whole-body visualization
✅ Dynamic
❌ Limited spatial resolution : tenths of a millimeter
❌ Acquisition time
❌ Expensive
Doppler Ultrasound imaging
✅ Real time
✅ Blood flow assessment
✅ No radiation
❌ Low resolution
❌ Mainly 2D
Ultrasound Localization Microscopy

## Slide 7
Some Basics before ULM
●
Point Spread Function
(PSF)
B-Mode Image

## Slide 8
●
●
B-Mode Image
2 PSFs are separable
Some Basics before ULM

## Slide 9
●
●
B-Mode Image
2 PSFs are Undistinguishable
Diffraction limit
Some Basics before ULM

## Slide 10
B-Mode Image
That is speckle
Interference of scattered waves
Some Basics before ULM

## Slide 11
Ultrasound Localization Microscopy
1) experimental setup
2) ULM processing pipeline
3) ULM Maps

## Slide 12
Ultrasound Localization Microscopy
1) experimental setup
2) ULM processing pipeline
3) ULM Maps
Why microbubbles?
✅ They stay in the blood pool
✅ Very strong ultrasound scatterers
✅ Tiny size (~1–5 µm radius) : behave like point scatterers at MHz frequencies
✅ Gas-filled core : huge acoustic impedance mismatch with Blood/tissue
Acoustic impedance:
ρ = density, c = speed of sound
ρ ≈ 1000 kg/m³, c ≈ 1540 m/s
​≈1.5MRayl
Reflection coefficient
Bubble–tissue interface : ∣R∣≈1
≈99.9% of the intensity is reflected
Bubbles look extremely bright in ultrasound images.
Soft tissue :
ρ ≈ 1.2 kg/m³, c ≈ 340 m/s
​≈0.0004MRayl
Bubbles :

## Slide 13
Ultrasound Localization Microscopy
Errico et al,
J. Foiret et al,
Noémi Renaudin et al,
Alexandre Corazza et al,
Baptiste HEILES et al,
❌ Used mainly in small animals
❌ Typically 2D or small 3D volumes

## Slide 14
Ultrasound Localization Microscopy
Errico et al,
J. Foiret et al,
Noémi Renaudin et al,
Alexandre Corazza et al,
Baptiste HEILES et al,
❌ Used mainly in small animals
❌ Typically 2D or small 3D volumes
✅ Map whole-organ microcirculation in 3D at high resolution
✅ 100 × 100 × 100 mm³ @ 100 µm resolution @ 300 Hz ❌ Impossible with traditional probes
✅ Multi-lens array + 3D ULM
✅ Up to 120 × 100 × 82 mm³ volume
✅ 125–200 µm resolution at 312 Hz
✅ Low-cost, scalable system
Scientific approach

## Slide 15
3D ultrasound imaging
2D Matrix Arrays
❌ Respect pitch=half wavelength
❌ Small Transducer elements
❌ Large number of elements (>1024)
❌ Complex connectivity
❌ Cable size
❌ Data transfer/storing problem
Example :
Consider a 128 element linear array (38mm aperture)
The same array in 2D have 128x128 element
64 Verasonics systems are needed
Paper needs : around 100mm wide aperture

## Slide 16
3D ultrasound imaging
2D Matrix Arrays
❌ Respect pitch=half wavelength
❌ Small Transducer elements
❌ Large number of elements (>1024)
❌ Complex connectivity
❌ Cable size
❌ Data transfer/storing problem
Example :
Consider a 128 element linear array (38mm aperture)
The same array in 2D have 128x128 element
64 Verasonics systems are needed
Paper needs : around 100mm wide aperture
Channel multiplexing
RCA Arrays
Micro-Beamforming
Sparse Arrays
GE 4Vc-D XDclear
We can still do 2D matrix array
Not multiplex & not micro-beamformed
Low number of elements
Large elements

## Slide 17
3D ultrasound imaging
Small Elements
Large Elements
✅ Good divergence
❌ High Impedance
✅ Low Impedance
❌ Highly Directive

## Slide 18
3D ultrasound imaging
Small Elements
Large Elements
✅ Good divergence
❌ High Impedance
✅ Low Impedance
❌ Highly Directive
Element width = λ
Element width = 3λ

## Slide 19
3D ultrasound imaging
Small Elements
Large Elements
✅ Good divergence
❌ High Impedance
✅ Low Impedance
❌ Highly Directive
Simplified electrical circuit
including small element
Element capacitance :
Element width = λ
Element width = 3λ
Few Elements + Large Elements + Diverging Elements
The winning combination

## Slide 20
Making Large Elements Diverging
Snell’s low
Refraction
c1 > c2 will results in
c2 > c1 will results in
✅ We can use this to create a lens to diverge the acoustic wave of a large element

## Slide 21
Snell’s low
Refraction
c1 > c2 will results in
c2 > c1 will results in
✅ We can use this to create a lens to diverge the acoustic wave of a large element
Plano-convex lens (c1 = 2570 m/s) combined to plano-concave lens (c2 = 1015 m/s)
Making Large Elements Diverging

## Slide 22
Element Directivity
✅ 3D maps of maximum transmitted acoustic pressure field
✅ Volume of 60 × 60 × 60 mm³
✅ The size of an element was 4.5 × 4.5 mm² representing 3 times the acoustic wavelength
✅ Acoustic lens increase divergence
✅ The -6db directivity : large element with lens (33.17°), the large element (10.27°) and the small element (51.13°).
✅ Large element increase the transmitted energy

## Slide 23
The Multi-Lens Array
Aperture size of the probe was 104 × 82 mm²
Central frequency of 1MHz (60% BW)
Three acoustic principal layers :
252 large transducer matrix piezocomposite elements (14 × 18, 4.5 × 4.5 mm, pitch 3.65λ)
Diverging acoustic lenses (plano-convex lenses), also arranged in 14 × 18 configuration.
Flat lens layer (silicon plano-concave lens) was molded above the diverging lens layer

## Slide 24
Array sensitivity
Dense array ( element size : λ/2 × λ/2, aperture size : 14 × 11 mm)
Sparse array : no information given
Sensitivity (105 × 100 × 100 mm volume):

## Slide 25
Sparse Array or Not ?
The beampattern
is the Fourier transform of
the aperture function
2 important metrics
Their pitch is d=3
Whether they like it or not :
their probe is sparse
Main-lobe width : affect resolution
Grating lobe : affect contrast
Extreme case :
Pitch condition

## Slide 26
Imaging sequence
Diverging wave emissions : short pulses of two cycles (1 MHz).
PRF = 5000 Hz  312.5 Hz
Tube phantom: 1000 blocks
Ex vivo heart: 400 blocks
In vivo kidney: 290 blocks
In vivo liver: 133 blocks
Total acquisition time
Tube Phantom: 13 min
Heart: 5 min
Kidney: 5 min
Liver: 162 s (~2.7 min)
Data acquired in blocks
In vitro & ex vivo: 830 ms → 256 volumes/block
In vivo: 1.2 s blocks
Blocks spaced every 300 ms (for data saving)
Number of blocks per experiment

## Slide 27
Results : PSFs
PSFs at depths of 3 cm, 4.5 cm, and 6 cm
252 sources :
Grating lobe level of approximately –25.50 dB, –24.57 dB, and –25.36 dB
Lateral resolution was found to be 1.15 mm, 1.17 mm, and 1.26 mm
16 sources
Probably around -17dB
Probably around 1.20mm
Simulation
Experimental
Tip of a needs
16 sources :
Grating lobes of -22.40 dB and
lateral resolution of 1.06 mm

## Slide 28
Reconstructed MB density volume
large imaging volume of 100 × 93 × 30 mm³
No grating lobes : filtered by tracking algorithm
Diameter was measured : 916 ± 38 μm
Agreement with the tube diameter (870 μm).
Measured flow in a agreement with the imposed flow
Polyethylene tube immersed in a water.
Diameter 870 μm and depth = 60 mm.
100 μL of microbubbles mixed with 100 ml injected
Rate of 5 mL/h , 100 mL/h , 150 mL/h , and 200 mL/h
Results : Flow tube

## Slide 29
MB density map of the coronary network
3D flows velocity map
Experiment on the isolated perfused heart
isolated perfused heart
Normal heart flow : Left ventricle → Aorta → Coronary arteries → Heart muscle
Retrograde :
pump fluid backward into the aorta
pressure forces the fluid into the coronary arteries
Oxygenated Krebs-Henseleit solution
Perfusion rate: 0.25 L/min=250 mL
Heart motion was reduced by setting the solution to 3°C : slow metabolism and contractions
Volume: 120 × 100 × 82 mm³
Full coronary network visualized
Flow velocity: 10 – 300 mm/s
Vessel radius: 75 – 600 µm
Results : Isolated Heart

## Slide 30
3 swines fully anesthetized and mechanically ventilated
Real-time B-mode with Philips® Lumify (S4-1U probe) : for probe guidance
2D probe fixed using an articulated arm
Microbubbles :1mL(0.5 mL every minute) Intravenous injection (ear vein)
Acquisitions triggered between respiratory phases (ECG)
In vivo porcine liver Setup
In vivo porcine kidney Setup
Results : In-vivo

## Slide 31
MB density map of the kidney network
3D flows velocity map
Volume: 65 × 100 × 82 mm³
Velocity: up to 300 mm/s
More complex : harder due to organs/ribs/motion
Volume: 60 × 80 × 40 mm³
Vessel radius: 70 – 400 µm
MB density map of the Liver network
3D flows velocity map
Results : In-vivo

## Slide 32
Proposed a multi-lens array + 3D ULM to image whole-organ microvasculature
Validated in simulations, phantoms, ex vivo heart, in vivo kidney & liver
Achieved: 125 µm spatial resolution 312 Hz temporal resolution
Cost of the system less than €200k
Spatial resolution, estimated at 75×75×75 μm not enough to visualize small capillaries (10 μm).
Improved by developing multi-lens arrays operating at higher frequency (2 to 3 MHz)
Detection based on correlation using gaussian modeled PSF
Low quality B-Mode due to grating lobes
The size of the array may cause patient discomfort
Conclusions
Limitations

## Slide 33
The media attention